Robust Intracardiac Communication under

Pathological Dynamics: Beat-Synchronous 2-DOF

Gain Compensation for Leadless Pacemakers

Dongming Li, Student Member, IEEE,

Abstract—Implantable  Medical  Devices  (IMDs)  are  evolving into collaborative networks, demanding robust, ultra-low-power communication  within  dynamic  and  often  pathological  biolog- ical  environments.  Multi-chamber  leadless  pacemakers  (LCPs) present the ultimate test for this robustness, as the intracardiac channel  exhibits  not  only  periodic  fluctuations  during  healthy rhythms but also unpredictable, aperiodic fades under patholog- ical conditions like Atrioventricular (AV) block. To address this challenge,  an  adaptive  receiver  based  on  a  Super-Regenerative Receiver  (SRR)  and  a  hybrid   analog-digital  Automatic   Gain Control   (AGC)   system  is   proposed.  At  its   core  is   a  beat- synchronous,  two-degree-of-freedom  (2-DOF)  control  strategy that  uses  the  peak  RSSI  of  each  beat  as  an  ECG-free  proxy for  the  cardiac  rhythm,  enabling  active  signal  stabilization  via predictive  feedforward  and  incremental  Proportional-Integral- Derivative  (PID)  control.  The  architecture  was  validated  using a  dual-path  evaluation  methodology:  a  programmable  ex-vivo porcine heart platform was developed to emulate the dynamics of Mobitz  Type  I Atrioventricular Block  (AVB) for  end-to-end link assessment, complemented by Hardware-in-the-Loop (HIL) simulations for transient characterization. Experimental results demonstrate that the proposed 2-DOF controller can stabilize the input signal to within 1%  of its target amidst channel variations exceeding 15 dB, translating into a Bit Error Rate (BER) reduc- tion  of one  to  two  orders  of magnitude,  with  an  improvement factor exceeding 60-fold at 5 kbps. This design paradigm provides critical  support  for  reliable  communication  in  multi-chamber LCPs  under  pathological  conditions  and  is  extensible  to  other networked IMDs operating in dynamic biological environments.

Index  Terms—Leadless  Pacemakers  (LCPs),  Galvanic   Con- ductive Communication (GCC),two-degree-of-freedom control (2- DOF)  Automatic  Gain  Control  (AGC),  Super-regenerative  Re- ceiver  (SRR), Pathological  Channel Dynamics, Atrioventricular Block.

# I.  INTRODUCTION

I

MPLANTABLE Medical Devices (IMDs) are undergoing a  paradigm  shift  from  isolated,  standalone  units  to  dis- tributed, networked systems  [1]–[4]. The future of advanced healthcare  will  rely  on  deploying  multiple  miniature,  intel- ligent  sensing,  monitoring,  or  therapeutic  nodes  within  the human body, forming cooperative  “In-Body Networks”  [5]– [7]. To realize this vision, a central technological bottleneck must be overcome: establishing a long-term, stable, and ultra- low-power wireless communication link within the complex and perpetually dynamic biological environment.

To enable cooperation among multi-node IMDs, academia has explored various wireless technologies. However, conven- tional radio-frequency (RF) communication faces severe signal attenuation within the human body, typically necessitating high

Fig.  1.   Illustration of the cardiac cycle, correlating the key electrical events of the electrocardiogram (ECG) with their corresponding mechanical phases, such as atrial systole and ventricular ejection.

power  consumption  to  maintain  a  link  [8].  While  magnetic inductive coupling offers lower power, it is constrained by its short communication range  and  stringent  alignment require- ments [9]. Ultrasonic communication, in turn, is hampered by its inherent directionality, making it unreliable for non-line-of- sight scenarios [10]. These intrinsic limitations severely curtail their  suitability for the next generation of networked IMDs, which demand miniaturization and ultra-low-power operation. In this context, Galvanic Conductive Communication (GCC) emerges as a promising and potentially disruptive technology. GCC cleverly utilizes human tissue itself as the transmission medium,  propagating  information  via  a  faint  electric  field. This approach not only offers the potential for extremely low power  consumption  and  simpler  circuit  implementation  but also  provides  enhanced  security,  as  the  signal  is  naturally confined  within  the  body.  Given  these  distinct  advantages, GCC has become one of the most promising candidates for enabling next-generation networked IMDs [11]–[14].

Among all implantable communication scenarios, the multi- chamber Leadless Pacemakers (LCPs)  system represents the ultimate  “proving  ground” for the robustness  of GCC tech- nology  [15]–[17]. Consequently, extensive research has been dedicated  to  characterizing  the  Intracardiac  Communication (CIC)  channel. Early investigations predominantly relied  on static or quasi-static models, employing finite element anal- ysis,  numerical  simulations,  and  in-vivo  animal  studies  to delineate fundamental transmission properties [18]–[21]. More recently, as research has matured, the dynamic nature of the CIC  has  gained  attention.  For  instance,  the  work  of  [22] provided valuable experimental insights into normal cardiac cycle  dynamics  via  an  ex-vivo  platform,  while  our  group’s

Fig. 2.  Schematic diagram of the cardiac conduction system and a represen-

tative Atrioventricular Block (Second-degree Type I).

prior work established a time-varying equivalent circuit model for such dynamic channels [23].

These  foundational   studies  are  indispensable  first   steps.  However, a critical challenge emerges as research transitions  from theoretical feasibility to clinical robustness. Many exist-  ing CIC transceiver designs, in their early stages, inevitably  rely on static or averaged channel models [20], [24], [25]. This  idealized approximation, while convenient for initial analysis,  masks a crucial fact: the intracardiac channel dynamics exist  on two distinct levels.The first level is the “healthy dynamic.” As  illustrated  in  Fig. 1,  the  regular  heartbeat  induces  pre-  dictable, periodic fluctuations in channel gain. This baseline  dynamic  itself precludes  the  use  of static  receivers  and  es-  tablishes the fundamental necessity for employing automatic  gain  control  (AGC).  However,  a  second,  more  severe  and  clinically  significant  level  is  the  “pathological  dynamic.”  A  representative form of this condition, an Atrioventricular (AV)  block as shown in Fig. 2, is characterized by intermittent con-  duction failure that breaks the original predictable rhythm and  introduces irregular, abrupt, and deeper channel fades. Under  these  conditions,  a  traditional  feedback  control  mechanism,  which relies purely  on  a  “post-event” reaction,  struggles  to  maintain link reliability in a timely and stable manner due to  its inherent response latency, forming a fundamental barrier to  translating prototypes into clinically reliable devices.

The  choice  of  receiver   architecture  becomes  paramount in  addressing  this  ultimate  challenge.  Given  the  extremely stringent constraints on power and complexity in LCPs, the Super-Regenerative  Receiver  (SRR)  architecture  emerges  as one of the most promising candidates, by virtue of its inherent simplicity and immense potential for achieving high sensitivity at very low power [26], [27]. The SRR architecture, however, is  a  double-edged  sword:  its  performance  exhibits  extreme sensitivity to variations in the input signal amplitude [28], [29]. While this characteristic can be exploited for high gain under stable channel conditions, it transforms into a critical liability in  the  dynamic  CIC  environment  we  face,  necessitating  a sophisticated compensatory mechanism to unlock the SRR’s full potential.

Therefore, AGC is an indispensable component for ensur- ing  system robustness. Yet, conventional AGC  solutions  are not directly applicable to LCPs. Fully analog AGC systems, while offering rapid response, often suffer from poor linearity, imprecise  gain  control,  and  a  fixed-threshold  operation  that

limits  adaptability  in  complex,  varying  environments   [30]. Conversely,  fully  digital  AGC  solutions,  while  capable  of high precision and flexibility, come at the cost of significant computational  overhead  and  power  consumption,  which  is incompatible with the microwatt-level budget of LCPs  [31]. Therefore, there is an urgent need for an innovative solution that can achieve a delicate balance among speed, accuracy, and low power  consumption,  and possess  the  ability  to  actively preset rhythmic channel hopping.

In response to this imperative, this paper introduces a novel receiver front-end based on an SRR, deeply co-designed with a  hybrid  analog-digital  AGC  architecture.  This  architecture strategically combines the rapid response of an analog front- end  with  the  fine-grained  control  of  the  digital  domain  to achieve robust compensation for dynamic CIC channel gain fluctuations.  The  main  contributions  of  this   work  are  as follows:

•   A  beat-synchronous  two-degree-of-freedom  control  (2- DOF)  active  compensation  strategy  is proposed,  which operates  without  an  external  ECG.  The  peak  RSSI  of each  beat  is  used  as  a  proxy  for  the  cardiac  rhythm, combining predictive feedforward with incremental PID control to proactively stabilize the gain against aperiodic, pathologically-triggered step changes.

•   A  programmable  ex-vivo  platform  was  developed  to emulate  the  pathological  dynamics  of  Mobitz  Type  I Atrioventricular Block (AVB). This controllable and re- peatable  setup enables the  systematic evaluation  of the communication link under clinically relevant and repro- ducible adverse channel conditions.

•   A dual-path quantitative evaluation framework was estab- lished.  By  integrating  Hardware-in-the-Loop  (HIL)  for transient  characterization  with  ex-vivo  experiments  for end-to-end communication assessment, a clear causal link from  control  metrics  to  link  reliability  was  built  and validated across multiple heart samples.

# II.  METHODS

## A.   Challenges  of  the Dynamic Intracardiac  Channel and Its Impact on SRR Performance

The  fundamental  challenge  for  efficient  communication between  multi-chamber  LCPs  arises  from  the  dynamic  and complex physiological environment of the heart. The periodic electromechanical activity of the heart induces significant and time-varying fluctuations in the CIC channel gain, Gc (t) [23]. Under pathological conditions requiring pacing intervention, such as “missed beats” caused by conduction block, the chan- nel dynamics become even more complex and unpredictable than during a normal sinus rhythm. This study, therefore, aims to design an adaptive receiver capable of maintaining robust communication in such a harsh and non-stationary channel.

Given the ultra-low-power constraints of LCPs, the SRR is a favored architecture due to its high  sensitivity and  simple structure.  However,  the  performance  of  an  SRR  is  excep- tionally sensitive to the input signal amplitude, Ain (t). This characteristic  stems  from  its  operating principle:  during  the regeneration phase, the net conductance of the oscillator tank

Preamp Preamp (AD8421)

| Hybrid-AGC VGA 2-DOF RSSI Controller |  | SRR Vref LC-Based SRO Envelope Detector | + - |
| --- | --- | --- | --- |
|  | Self-quench signal |  |  |

Fig.  3.   Block  diagram  of the proposed  adaptive receiver architecture. The core Hybrid-AGC module forms a closed loop, using a RSSI and a 2-DOF controller  to  dynamically  adjust  the  VGA,  thereby  delivering  a  stabilized signal to the SRR stage.

is negative (G = Go —GN  < 0), causing the voltage envelope Venv (t) to grow exponentially from an initial value set by the input signal Ain (t):

where Tregen  is the regeneration period. Consequently, within a fixed quench period, the final detected envelope amplitude is a monotonically increasing function of Ain (t). In a front-end without adaptive control, the SRR’s input amplitude Ain (t) is directly coupled with the time-varying channel:

Ain (t) = Atx · GLNA · Gc (t)                    (2)

where Atx and GLNA  are the transmitted signal amplitude and the low-noise amplifier (LNA) gain, respectively. Fluctuations in  channel  gain  Gc (t)  directly  translate  into  variations  in Ain (t), which in turn cause drastic changes in the effective signal-to-noise ratio (SNReff ) at the SRR’s decision point. This relationship can be modeled as a monotonic function:

Pe (t) = fBER (SNReff (t)), SNReff (t) Ⅸ An (t)    (3)

where fBER (·) is a monotonically decreasing function. During deep channel fades, a  sharp drop in  SNReff (t) causes the bit error rate Pe (t) to approach 0.5, leading to instantaneous link failure. Therefore, stabilizing Ain (t) is a prerequisite for reliable communication. Notably, the control scheme proposed herein does not rely on an external electrocardiogram (ECG) signal;  when  the  ECG  is  unavailable,  the  system  can  au- tonomously  extract rhythmic information from the  envelope of the received signal strength indicator (RSSI).

## B.   Adaptive Receiver Architecture with Hybrid Gain Control

To  address  the  aforementioned  challenges,  we  propose an  adaptive  receiver  architecture  centered  around  a  hybrid analog-digital automatic gain control (Hybrid-AGC) loop, as illustrated  in  Fig. 3.  The  core  objective  of this  architecture is to dynamically adjust the total front-end gain to precisely compensate for fluctuations in the channel gain Gc (t), thereby providing a constant-amplitude input signal to the subsequent SRR core.

As depicted in the diagram, the receiver chain is divided into three functional stages. The signal first enters a fixed pre- amplification  stage  (Preamp),  followed  by  the  Hybrid-AGC module,  which  acts  as  the  core  regulation  unit.  This  mod- ule forms a sophisticated negative feedback loop: a variable

| Algorithm 1 Beat-Synchronous with RSSI Prediction | 2-DOF | Gain | Stabilization |
| --- | --- | --- | --- |
| Input: Target level xig , Kf , α, PID gains (Kp , Ki , Kd , Kaw ), bounds umin , umax , slew limit ∆umax Output: Vc [k] 1: Initialize: Iprev ← 0, Dprev ← 0, eprev ← 0 2: prev ← xig , ufinal prev ← 0, uPID prev ← 0, uff ← 0 3: while system running do 4: x[k] ← Sample and digitize RSSI dB 5: if new heartbeat detected then 6: xp ← Peak RSSI of previous beat dB 7: ← (1 - α)prev + αxp 8: uff ← Kf (xig - ) 9: prev ← 10: end if :: 13: uPID [k] ← uPID prev + Kp (e[k] - eprev ) + KiTse[k] + (D[k] - Dprev ) ▷ incremental PID 14: utot [k] ← uff + uPID [k] 15: ufinal [k] ← clip(utot [k], ufinalprev ± ∆umax ) 16: ▷ slew limit 17: usat [k] ← clip(ufinal [k], umin , umax ) 18: ▷ saturation 19: Vc [k] ← DAC(usat [k]) 20: Update history for next iteration: 21: Dprev ← D[k], eprev ← e[k] 22: uPID prev ← uPID [k], ufinal prev ← ufinal [k] 23: end while |  |  |  |

gain amplifier  (VGA) amplifies the  signal, and its output is simultaneously routed  to  the  final  SRR  detection  stage  and monitored  in  real-time  by  a  RSSI.  The  RSSI  converts  the signal envelope’s  strength into  a corresponding DC voltage, which is then sampled by an analog-to-digital converter and fed to a digital 2-DOF controller. This controller compares the sampled value against a preset target, computes the error, and generates an updated gain control signal based on its control algorithm. This digital signal is converted back to an analog voltage by a digital-to-analog converter to precisely regulate the  VGA’s  gain.  This  completes  the  closed-loop regulation, ensuring that the signal strength delivered to the SRR remains stable regardless of input channel variations.

## C.  Loop Modeling, Stability, and Implementation

To ensure the closed-loop system exhibits robust and pre- dictable  dynamic  performance,  we  first  establish  a  control- theoretic model for the AGC–SRR system, which guides the controller’s design and tuning. We approximate the entire path from the VGA gain control input to the ADC output in the logarithmic  domain  as  a  first-order  continuous-time  system with  a  pure  delay  τd:  x.(t)  =  —(1/τ)x(t) + γcu(t — τd ). Discretization using a zero-order hold yields its discrete-time equivalent model:

x[k + 1] = ϕx[k] + γu[k — d] + w[k]              (4)

where x[k] is the sampled RSSI value, u[k] is the VGA gain control signal, and the mapping from continuous to discrete

parameters is:

ϕ = e −Ts /τ ,    γ = γc τ(1 − e−Ts /τ ),    d = ⌈τd /Ts⌉     (5)

The term w[k] represents model uncertainty and noise. The transfer function of the plant, P(z), is therefore:

The   system’s   open-loop   transfer   function   is   L(z)   = C(z)P(z).  Our  design  objective  is  to  maintain  closed-loop stability under all operating conditions, with a specific target of a phase margin ≥ 45◦ . This provides a clear optimization goal  for  tuning  the  PID  parameters  (Kp , Ki , Kd ).  The  ap- proximate relationship between the settling time tsettle  and the controller parameters further reveals their impact on dynamic performance:

To translate this theoretical model into a high-performance practical controller, we integrated several key implementation techniques.  To  suppress the  amplification  of high-frequency noise by the derivative term, we employed a first-order low- pass filter to smooth the rate of error change, with its update rule given below (recommended parameters: Td  = 2Ts , N = 10):

Simultaneously, to prevent integral windup when the VGA gain saturates, we implemented an anti-windup mechanism via back-calculation, where the integral accumulator is corrected based on the saturation error:

I[k + 1] = I[k] + KiTse[k] + Kaw (usat [k] − u[k])     (9)

where  usat [k]  is  the  saturated  control  output.  Furthermore, during the preamble of a data packet, we temporarily freeze the integral and derivative channels, relying solely on proportional control for rapid coarse adjustment to avoid overshoot from the initial step response.

## D.  Beat-Synchronous  2-DOF  Control  Algorithm  with  RSSI Prediction

Despite these optimizations, a conventional feedback-only PID controller is still limited by its loop bandwidth in terms of response speed and overshoot when facing the abrupt, step- like  changes  in  channel  gain  inherent  to  the  cardiac  cycle. While a well-tuned PID can manage the rhythmic, predictable steps  of a healthy heart to  some  degree, its purely reactive nature proves critically insufficient for the chaotic dynamics of pathological conditions. The core challenge shifts from man- aging periodicity to compensating for unpredictability during events like the  “missed beats” in  an Atrioventricular block. To overcome this bottleneck of unpredictability, we propose a  2-DOF  control   strategy  that  combines  beat-synchronous feedforward control with PID feedback.

The fundamental principle of this strategy is rooted in the direct mapping between cardiac mechanics and received signal

strength. As established in prior work  [17],  [32], the heart’s mechanical  pumping  action,  characterized  by  the  cyclical filling  and  ejection  of blood—causes  significant  changes  in both the geometry between the electrodes and the electrical properties of the transmission medium. These hemodynamic shifts directly modulate the channel gain (Gc (t)). The RSSI provides a real-time measurement of this gain, effectively serv- ing as an electrical proxy for the heart’s mechanical rhythm. It is this robust link that allows our system to leverage the heart’s rhythmicity. By analyzing the RSSI envelope, the controller can  autonomously identify beat boundaries  and characterize each beat’s channel quality by its peak RSSI value forming the basis for prediction.

Based  on  this,  we  employ  a  computationally  lightweight Exponential Moving Average (EMA) model to predict the peak RSSI of the next beat:

k+1  = (1 − α)k  + αxk ,    α ∈ [0.2, 0.3]        (10)

where k+1  is the prediction  for the next beat’s peak RSSI.

Our 2-DOF controller combines predictive compensation for the  slowly  varying  component  of  the   channel  (beat-level jumps)  with  feedback  regulation  for  fast-varying  residuals. Ideally,  the  required  VGA  gain  compensation  (in  dB)  is

gga (t) = x* −(gc (t)+glna +gtx ). Upon detecting the start of

a new beat, the controller first calculates a feedforward control signal, uff :

uff [k + 1] = Kf (x* − k+1),    uff  ∈ [−8, 8] dB       (11)

This feedforward term is added to the incrementally imple- mented PID feedback output, uPID , to form the total control signal:

(12)

uPID [k] = uPID [k − 1] + Kp (e[k] − e[k − 1])

+ KiTse[k] + (D[k] − D[k − 1])

utotal [k] = uff [k] + uPID [k]                    (13)

To prevent rapid changes from stressing the hardware, we also apply a slew-rate limiter to the total control output:

(14)

ufinal [k] = clip(utotal [k], ufinal [k − 1]

− ∆umax , ufinal [k − 1] + ∆umax )

Furthermore,  to  enhance  robustness,  the  system  includes a de-rating  strategy: if the prediction error remains below a threshold for several consecutive beats, the feedforward gain Kf   is temporarily zeroed,  and the  system reverts to  a pure PID mode until the next large jump is detected. The complete control logic, which integrates prediction, feedforward, feed- back, and multiple robustness enhancements, is formalized in Algorithm 1.

# III.  EXPERIMENTAL METHODOLOGY AND SETUP

To rigorously validate the proposed adaptive receiver archi- tecture  and  its  core  2-DOF  control  strategy,  we  established a  high-fidelity  ex-vivo  experimental  methodology.  The  cor- nerstone  of this  methodology  is  an  isolated  heart perfusion platform in Fig. 4,  which utilizes  a freshly  excised porcine heart as its biological core to maximally replicate the authentic

Fig. 4.  Photograph of the assembled experimental testbed, showing the peri- staltic pump system, the cannulated isolated porcine heart, signal generation instruments, and the custom-developed receiver module. A 1:1 balun is used for current isolation.

TABLE I

CARDIAC CYCLE PHASES AND CORRESPONDING CHAMBER VOLUMES

| Phase | Duration (s) | Left Volume (mL) | Right Volume (mL) |
| --- | --- | --- | --- |
| Isovolumic Relaxation | 0.07 | 120.00 | 150.00 |
| Rapid Filling | 0.11 | 173.44 | 202.88 |
| Diastasis | 0.22 | 191.25 | 220.50 |
| Atrial Systole | 0.10 | 215.00 | 244.00 |
| Isovolumic Contraction | 0.05 | 215.00 | 244.00 |
| Rapid Ejection | 0.10 | 151.67 | 181.34 |
| Reduced Ejection | 0.15 | 120.00 | 150.00 |

intracardiac communication environment. A dual-channel peri- staltic pump system circulates a specially formulated blood- mimicking fluid, matched to the conductivity of human blood at 37°C [21], [33]. This setup, governed by a host computer executing  precise  temporal  control  over  chamber  volumes, achieves  a  high-fidelity  simulation  of  intracardiac  hemody- namics.

We contend that a truly robust adaptive system must main- tain performance under harsh channel conditions characterized by  both  gradual  and  abrupt  variations.  Consequently,  we deliberately selected and implemented Mobitz Type I AVB as a representative pathological channel model. The key advan- tage of this model is its inherent Wenckebach phenomenon, where channel parameters progressively deteriorate before  a “dropped beat,” and the dropped beat itself manifests as an abrupt  step-change  in  channel  gain.  Our  platform  was  pro- grammed to precisely replicate the hemodynamic  signatures of this AVB model, based on published human cardiac volume data  [22],  [34]–[37],  with  the  specific  parameters  for  each phase of the cardiac cycle detailwith the specific parameters for each phase of the cardiac cycle detailed in Table I. The dynamic mappings for both normal sinus rhythm (NSR) and

our AVB model are illustrated in Fig. 5and Fig. 6, respectively. 	Upon  this  high-fidelity  platform,   we  instituted  a  direct comparative  evaluation  framework.  This  framework  bench- marks our proposed 2-DOF controller, which integrates beat- synchronous predictive feedforward, against a well-tuned con- ventional feedback-only PID controller (serving as the base- line)  and  an open-loop configuration with no AGC.  System performance was quantified using a suite of key metrics cov- ering both transient dynamics and steady-state communication quality. The former includes settling time (tsettle ) and maximum overshoot, while the latter is assessed by the end-to-end Bit Error Rate (BER). Across all experiments, the physical link setup remained constant: a 1MHz OOK signal with an average power of —10dBm was generated and coupled to the right

atrium, while our custom-developed SRR receiver, integrating the different control strategies, was connected to an electrode in the right ventricle.

Our evaluation strategy was strategically bifurcated into two complementary phases. First, to precisely assess the intrinsic transient response capabilities of the controller, independent of the complexities of the biological specimen, we employed a HIL simulation strategy. In this setup, the controller hardware under  test  (MCU)  and  its  actuator  (VGA)  were  physical devices, but their RF environment was digitally simulated in real-time  by  a  host  computer.  The  computer  calculated  the requisite signal strength at the SRR input based on predefined channel gain  step  sequences  (5  dB,  10  dB,  15  dB)  and  the MCU’s  real-time  output,  then  converted  this  value  into  an analog RSSI signal via a high-speed DAC to close the loop. Analyzing the controller’s response to these step disturbances allowed  for  precise  quantification  of  its   settling  time  and maximum overshoot.

Second, communication reliability, the ultimate measure of the system’s practical utility, was evaluated under the contin- uously dynamic pathological channel conditions generated by the isolated heart platform. As depicted in Fig. 7, a complete transceiver and analysis chain was established to conduct long- duration transmissions  of a pseudo-random binary  sequence (PRBS)  for  the  no-AGC,  PID,   and  2-DOF  configurations under the simulated AVB rhythm, enabling the calculation of BER at various data rates. To ensure the scientific rigor and reproducibility  of all  findings, the  entire  set  of experiments was repeated in three independent porcine heart preparations (N=3),  with  five  technical  replicates  for  each  measurement (n=5).

# IV.  EXPERIMENTAL RESULTS AND ANALYSIS

A series of systematic ex-vivo experiments were conducted to  comprehensively  evaluate  the  proposed  adaptive  receiver architecture and its core 2-DOF control strategy. This section sequentially constructs a complete argumentative chain: first, by presenting the true characteristics of the dynamic intrac- ardiac  channel,  we  reveal  and  quantify  the  core  challenges of  this  research;  second,  we  demonstrate  the  fundamental necessity of a basic AGC loop as a countermeasure; finally, through rigorous comparison with a traditional PID baseline, we will demonstrate the superior performance of our proposed

Fig. 5.  Dynamic mapping between ECG and right-sided chamber volume for

the simulated NSR, serving as the baseline model for a healthy heart.

Fig. 6.  Pathological cardiac cycle simulation corresponding to our represen- tative Atrioventricular Block (AVB) model. The map illustrates an instance of intermittent conduction failure, where a P-wave is not followed by a QRS complex, leading to an irregular hemodynamic profile.

2-DOF  controller  in  achieving  robust  communication  from both transient response and statistical perspectives.

## A.   Characterization of Channel Dynamics under Healthy and Pathological Rhythms

Direct experimental characterization of the RA-RV intrac- ardiac communication channel, as depicted in Fig. 8, reveals its inherent dynamic properties, providing a clear rationale for the design of an  adaptive receiver. Under  simulated healthy heart conditions which is  shown Fig. 8(a), the channel path gain exhibits pronounced yet highly periodic fluctuations of up to  10–15 dB, tightly synchronized with the cardiac cycle. This baseline  dynamic makes it  clear that  even under ideal conditions, a static receiver cannot guarantee stable communi- cation quality, thus establishing the fundamental necessity for employing AGC.

However, the core challenge of this research is fully mani- fested when simulating pathological conditions. As shown in Fig. 8(b), under the Mobitz Type I AVB model, the predictable rhythm of the channel is completely disrupted. It is replaced by a harsh environment dominated by irregular, abrupt, and deep

Fig.  7.     Schematic   of  the  experimental  setup  for  BER   evaluation.  The configuration integrates a signal generator (SG), the ex-vivo heart platform, the adaptive SRR, and a data acquisition (DAQ) system.

signal fades. This qualitative shift from periodic fluctuations to unpredictable, event-driven transients imposes requirements on the adaptive system that are far beyond the conventional. It indicates that a traditional feedback controller, which merely reacts  to  errors,  will  inevitably   suffer  from  performance degradation when coping with such sudden fades. Therefore, to maintain robust communication link integrity under these extreme conditions, an advanced control  strategy capable of predictive and rapid response becomes essential.

## B.   Fundamental  Effectiveness  of  the AGC  System  in  Signal Stabilization

Before introducing the advanced 2-DOF controller, we first had  to  establish  the  fundamental  effectiveness  of  the  AGC closed-loop  itself as  a  solution.  The basic  capability  of the AGC loop in stabilizing the severe channel fluctuations caused

Fig. 8.   Dynamic path gain variation of the RA-RV intracardiac channel. (a) Under simulated healthy heart conditions. (b) Under simulated representative pathological  (Mobitz Type I AVB) conditions. Results from three  different ex-vivo porcine heart preparations are shown.

by heartbeats is illustrated in Fig. 9, which uses the simulated pathological ECG in Fig. 9(a) as a temporal reference.

The waveform in Fig. 9(b) reveals the inherent instability without AGC, where the  signal amplitude at the SRR input fluctuates drastically between  1.5  mV and 5.0 mV with the simulated heartbeats; this would directly translate into dete- riorated  demodulation  performance.  In  stark  contrast,  when the  AGC  loop  is  activated,  the  VGA’s  control  voltage  (Vg ) accurately tracks the channel variations in real-time and in an inverse manner, as depicted in the trace of Fig. 9(c). The direct result, visualized in the  stabilized waveform of Fig. 9(d), is that the SRR’s input signal amplitude is successfully clamped at the 100 mV target level, with peak-to-peak residual fluctu- ations suppressed to within 1%. This result provides decisive evidence for AGC as a necessary technique to compensate for the dynamic channel.

## C.   Comprehensive  Evaluation of  Transient Response Perfor- mance

To quantify the advantages of our proposed 2-DOF strategy over  a  traditional  PID  baseline,  we  evaluated  the  transient response of the AGC control algorithm under a challenging 15  dB  channel  step  event. The results,  depicted in Fig. 10, reveal a stark performance difference. Following a sharp drop in channel gain, the feedback-only PID controller shows sig- nificant undershoot, with its signal amplitude falling to 87.7% of  its  target,  representing  a  maximum  deviation  of  12.3%, and requiring  85.2 ms to re-settle. Conversely, thanks to its predictive feedforward path, the 2-DOF controller restricts the maximum  deviation  to  only  9.1%  and  maintains  the  signal above 90.9% of its target value. It also dramatically reduces the settling time to just 20.3 ms, which is a greater than four- fold, or 76%, improvement over the baseline.

To generalize this significant advantage from a representa- tive event to a statistically meaningful conclusion, we further performed an aggregate analysis of key dynamic performance metrics measured across repeated experiments under various

(a) ECG Signal

6

(b) Received Signal Amplitude without AGC

Amplitude (mV)

4

2

0

30

(c) VGA Control Voltage

Vg (mV)

25

20

15

10

120

(d) Received Signal Amplitude with AGC

Amplitude (mV)

110

100

90

80

0.0       0.4       0.8        1.2       1.6       2.0       2.4       2.8       3.2

Time (s)

Fig. 9.   Effectiveness  of the hybrid AGC  system in  stabilizing the receiver input  signal.  (a)  Reference  ECG  signal.  (b)  Received  signal  amplitude  at the  SRR  input  without  AGC.  (c)  Dynamic  VGA  control  voltage  (Vg).  (d) Stabilized signal amplitude at the SRR input with AGC activated.

disturbance  scenarios, with the results presented in Fig. 11. The  statistical results irrefutably confirm the comprehensive superiority of the 2-DOF controller in dynamic performance, especially in response speed.

The   statistical   results   for   settling   time,   presented   in Fig. 11(a),  highlight  a  compelling  advantage  of the  2-DOF architecture, which reduces the settling time by more than 4 times, a reduction of approximately 76%, across all tested sce- narios. This superior performance extends to the suppression of maximum deviation, where the data in Fig. 11(b)  shows a consistent improvement that reduces the deviation by about 20–26% in all scenarios. Thus, the combined evidence from the in-depth analysis of a single typical event and this broader cross-scenario statistical validation robustly proves that the 2- DOF  architecture  possesses  far  superior  speed  and  stability in  coping  with  channel  transients  compared  to  traditional feedback control.

tsettle= 85.2 ms

20.3 ms

Min 87.7% (87.7 mV) Min 90.9% (90.9 mV)

Fig.   10.     Transient  response  comparison  between  the  2-DOF  and  PID controllers to a representative 15 dB channel gain step. The normalized SRR input signal demonstrates the superior performance of the 2-DOF controller in terms of both settling time and undershoot suppression.

| Settling time (ms) | 150 125 100 75 50 25 0 | (a) Settling Time PID 2-DOF 5 dB 10 dB 15 dB | Max Deviation (%) | 20 15 10 5 0 | (b) Max Deviation PID 2-DOF 5 dB 10 dB 15 dB |
| --- | --- | --- | --- | --- | --- |

Fig. 11.   Quantitative statistical comparison of transient performance metrics under different channel step disturbances. (a) Settling time (tsettle) comparison. (b)  Maximum  deviation  (%)  comparison.  Data  points  represent  individual trials, while bars and error whiskers indicate the mean and 95% confidence interval,  respectively.  The  2-DOF  controller  significantly  outperforms  the baseline PID across all scenarios.

## D.   Final Enhancement in Communication Reliability

The  ultimate  value  of  superior   signal  stability  must  be manifested in an improvement in actual communication perfor- mance. Fig. 12 first visually demonstrates the AGC system’s assurance of signal integrity in the time domain. Even under a dynamic channel, the envelope of the received OOK signal is maintained at a constant level, ensuring that the subsequent demodulation circuit can flawlessly recover the original base- band signal.

We quantified the final communication performance by the Bit Error Rate (BER), with the results presented in Fig. 13. The experimental results provide compelling evidence for the decisive role  of adaptive  control.  Compared to the  configu- ration without AGC, both AGC strategies (PID and 2-DOF) substantially reduced the BER across all tested data rates, with the 2-DOF configuration achieving an average improvement of 98.4% at a 5 kbps data rate.

More  importantly,  the  data  further  reveals  the  significant superiority of our proposed 2-DOF controller over the conven- tional PID baseline. Across all data rates, the 2-DOF controller consistently achieved a lower BER. This advantage becomes dramatically  more  pronounced   as  the  data  rate   increases, highlighting  its  resilience  to  inter-symbol  interference.  For instance,  while  the  performance  improvement  factor  of  the

Fig. 12.   OOK signal waveforms at different stages of the system with AGC activated.  (a) Transmitted baseband  signal.  (b) Received  signal  at the  SRR input. (c) Demodulated baseband signal at the SRR output.

| BER | 1 10 - 2 10 - 3 10 - 4 10 - 5 10 | NoAGC PID 2-DOF 500 1k 2k 3k 4k 5k | 100x 10x 1x | Improvement Factor |
| --- | --- | --- | --- | --- |

Data Rate (bps)

Fig.  13.    BER  performance  comparison  for  No-AGC,  PID,   and  2-DOF configurations  across  various  data  rates  under  the  simulated  pathological condition.  The  plot  highlights  the  significant  BER  reduction  achieved  by AGC,  with  the  2-DOF  controller  demonstrating  superior performance  over the traditional PID. Error bars represent the  standard deviation across N=3 independent heart preparations.

2-DOF  controller  over  the  No-AGC  case  is  approximately 24-fold  at  500 bps,  this  factor  escalates  to  over  63-fold  at 5 kbps. At this highest tested data rate, the 2-DOF controller reduced the average BER from the PID’s 9.0 × 10-5  down to 1.5×10-5, representing an additional 83.3% error rate reduc- tion on top of the already-optimized PID. This trend unequiv- ocally  demonstrates  that  the  superior  transient  response  of the 2-DOF controller translates into exceptional robustness for high-speed communication. Therefore, the proposed adaptive receiver architecture provides a key technical foundation for ensuring high-reliability communication for LCPs in complex pathological environments.

TABLE II

PERFORMANCE COMPARISON WITH RELATED RECEIVER PROTOTYPES

| Metric | Bereuter et al. [25] | Ryser et al. [24] | Khaleghi et al. [20] | Yuk et al. [38] | This Work |
| --- | --- | --- | --- | --- | --- |
| Primary focus | In-vivo channel characterization | Modulation scheme analysis | Broadband channel modeling | Low-power integrated circuit | Pathological channel robustness |
| Measurement Type | In-vivo | Ex-vivo | Ex-vivo | In-vivo/Ex-vivo | Ex-vivo |
| Frequency [MHz] | 1 | 1.5 | 0.01-10 | DC-20 | 1 |
| Receiver Power [mW] | >465 | 152 | 230 | 0.062 | 268 |
| Technology | Discrete components | Discrete components | Discrete components | 110nm CMOS | Discrete components |
| NSR Dynamics | × | × | × | × | √ |
| Pathological Dynamics | × | × | × | × | √ |
| Adaptive Gain Compensation | × | × | × | × | √ |

# V.  DISCUSSION

This study was centered on the core challenge of robust in- tracardiac communication under pathological cardiac rhythms. Experimental results demonstrate that the proposed adaptive receiver maintains a stable signal amplitude and significantly reduces the bit error rate, even in the presence of abrupt and unpredictable channel gain jumps triggered by events such as Atrioventricular block. This highlights a key  shift in design philosophy: intracardiac communication systems must evolve from adapting to regular, healthy rhythms to actively tracking and  suppressing aperiodic, pathologically-driven events. Our beat-synchronous 2-DOF strategy achieves this by using the RSSI as an ECG-free proxy for the cardiac rhythm, enabling beat-by-beat prediction and feedforward presetting, which is complemented by an incremental PID to  suppress intra-beat residual errors. This “predict and correct” combination allows the system to maintain a stable demodulation boundary during sudden transients.

As summarized in Table II, this work’s  starting point and evaluation  scenario  are  more  aligned  with  clinical  reality compared to prior art [20], [24], [25], [38]. Previous research has largely focused  on channel characterization, modulation analysis, or broadband modeling, almost entirely under static or quasi-static assumptions. While excellent progress has been made in low-power integrated circuit implementation, system- atically addressing the communication robustness challenges posed by cardiac dynamics, especially pathological dynamics, remains an open problem. This work focuses on this “patho- logical channel robustness” gap by advancing the evaluation scenario from  static to dynamic, and further to pathological conditions. As Table II also makes clear, existing prototypes generally lack adaptive gain compensation for dynamic chan- nels. In contrast, the hybrid architecture and beat-synchronous prediction  proposed  here  are  designed  to  balance  response speed and steady-state performance in a highly time-varying channel, directly filling this critical technological gap.

The  proposed  method,  being  designed  for  event-driven aperiodic  transients,  is  expected  to  be  extensible  to  other scenarios such as dropped beats, intermittent conduction ab- normalities, and highly irregular intervals. It should be noted that  the  current  ex  vivo  platform  cannot  cover  all  in-vivo complexities, such as autonomic nervous regulation, long-term

tissue-electrode  interface  changes,  and  respiratory  coupling. The prototype was also implemented with discrete components to validate the control concept and architecture, rather than as a final implantable device.

From  a  clinical  perspective, reliable  intracardiac  commu- nication  is  fundamental  for  synchronized  pacing  and  inter- chamber  coordination in multi-chamber pacemaker  systems. In  settings  dominated  by  pathological  rhythms,  the  ability to  proactively  preset  the  gain  and  rapidly  stabilize  the  link helps ensure the accurate delivery of rhythm management and therapeutic sequences. Future work will focus on the CMOS integrated circuit implementation of this architecture, balanc- ing  low  latency  with  fixed-point  complexity  while  ensuring stability margins. This will be followed by longer-term large- animal in-vivo studies covering a broader range of pathological types  and more complex physiological conditions to further validate the method’s transferability and long-term stability.

# VI.  CONCLUSION

This paper addressed the critical challenge of maintaining robust  intracardiac  communication  amidst  the  unpredictable channel fades caused by pathological cardiac events. To over- come  the  limitations   of  conventional  reactive  control,  we proposed and validated a novel adaptive receiver architecture featuring a beat-synchronous, 2-DOF Hybrid-AGC controller that proactively compensates for channel variations by using the RSSI as an ECG-free proxy for the cardiac rhythm. Sys- tematic evaluation demonstrated that the controller stabilizes the received signal to within 1% of its target despite channel gain fluctuations exceeding 15 dB. This stabilization translates into  a  bit  error  rate  improvement  of  up  to  98.4%,  which corresponds  to  a  reliability  enhancement  factor  of over  60. This  research,  therefore,  delivers  not  only  a  key  enabling technology for clinically reliable LCP systems but also offers a rigorously validated and generalizable design paradigm for the next generation of networked IMDs intended to operate robustly in harsh biological environments.

# REFERENCES

[1]  J. E. Ferguson and D. Redish, “Wireless communication with implanted medical  devices using the  conductive properties  of the body,” Expert Review of Medical Devices, vol. 8, pp. 427 – 433, 2011.

[2]  E. Zrenner,  “Will retinal implants restore vision?,” Science, vol.  295, no. 5557, pp. 1022–1025, 2002.

[3]  P. Laizou, “Signal-processing techniques for cochlear implants,” IEEE Engineering in Medicine and Biology magazine, vol.  18, no. 3, pp. 34– 46, 1999.

[4]  W. Germanovix and C. Toumazou, “Design of a micropower current- mode  log-domain  analog  cochlear  implant,”  IEEE  Transactions  on Circuits and Systems Ii: Analog and Digital Signal Processing, vol. 47, pp. 1023–1046, 2000.

[5]  M. Estudillo, D. Naranjo, L. M. Roa, and J. Reina-Tosina, “Intrabody communications (ibc) as an alternative proposal for biomedical wearable systems,” in  Handbook of Research on Developments  in E-Health and Telemedicine:  Technological  and  Social  Perspectives,  pp.   1–28,  IGI Global Scientific Publishing, 2010.

[6]  S.  Warren  and  B.  Natarajan,  “Wireless  communication  technologies for  wearable  systems,”  in   Wearable  Monitoring  Systems,  pp.  51–80, Springer, 2010.

[7]  Y.-S. Lin, C.-C. Wang, Y.-C. Liao, and S.-S. Lu, “Design and imple- mentation of intrabody communication hub/alarm unit in ibc platform for fall prevention system,”  Microwave and Optical Technology Letters, vol. 56, no. 10, pp. 2345–2351, 2014.

[8]  P. Bose, A. Khaleghi, M. Albatat, J. Bergsland, and I. Balasingham, “Rf channel  modeling  for  implant-to-implant  communication  and  implant to  subcutaneous  implant  communication  for  future  leadless  cardiac pacemakers,”  IEEE  Transactions on Biomedical Engineering, vol. 65, no. 12, pp. 2798–2807, 2018.

[9]  W.  Li,  Y.  Duan,  and  J.  M.  Rabaey,  “A  200mb/s  inductively  coupled wireless transcranial transceiver achieving 5e-11 ber and 1.5 pj/b trans- mit energy efficiency,” in  2018 IEEE International Solid-State Circuits Conference-(ISSCC), pp. 290–292, IEEE, 2018.

[10]  M. Seyedi, B. Kibret, D. T. Lai, and M. Faulkner, “A survey on intrabody communications for body area network applications,” IEEE Transactions on Biomedical Engineering, vol. 60, no. 8, pp. 2067–2079, 2013.

[11]  J.  Wu,   Y.  Liu,  Y.  Gao,  and  M.  Du,  “Equivalent  heart  model  for intracardiac  communication  channel  analysis,”  in   2020   Cross  Strait Radio Science & Wireless Technology Conference (CSRSWTC), pp. 1–3, IEEE, 2020.

[12]  L.  H.  Opie,   Heart  physiology: from  cell  to  circulation.     Lippincott Williams & Wilkins, 2004.

[13]  Wegmueller,  Marc,  Simon,  Oberle,  Michael,  Felber,  Norbert,  Kuster, Niels,  and F.  and,  “Signal transmission by galvanic coupling through the human body.,”  IEEE Transactions on Instrumentation  Measurement, 2010.

[14]  A.  K.  Teshome,  B.  Kibret,  and  D.  T.  H.  Lai,  “Galvanically  coupled intrabody  communications  for  medical  implants:  A  unified  analytic model,”  IEEE  Transactions  on Antennas   Propagation,  vol. 64, no. 7, pp. 2989–3002, 2016.

[15]  T. G. Middour, J. H. Chen, and M. F. El-Chami, “Leadless pacemakers: a review of current data and future directions,” Progress in Cardiovascular Diseases, vol. 66, pp. 61–69, 2021.

[16]  K. Malaczynska-Rajpold, M. Elliot, N. Wijesuriya, V. Mehta, T. Wong,  C. A. Rinaldi, and J. M. Behar, “Leadless cardiac pacing: new horizons,” Cardiology and Therapy, vol. 12, no. 1, pp. 21–33, 2023.

[17]  D. Li, Z. Wang, H. Wang,  S. H. Pun, P. U. Mak, A. Zhang, Y. Liu, H.-C. Li, Y. Gao, M. Du, et al., “Dynamic path gain compensation for enhancing intracardiac  communication in  leadless pacemakers,” IEEE Transactions on Instrumentation and Measurement, 2025.

[18]  L. Bereuter, T. Kuenzle, T. Niederhauser, M. Kucera, D. Obrist, T. Re- ichlin, H. Tanner, and A. Haeberlin, “Fundamental characterization of conductive intracardiac communication for leadless multisite pacemaker systems,”  IEEE transactions on biomedical circuits and systems, vol. 13, no. 1, pp. 237–247, 2018.

[19]  H. Wang, D. Li, M. I. Vai, S. H. Pun, J. Yang, H. C. Li, and Y. Gao, “Analysis and circuit design of imbalanced impedance channels for con- ductive intracardiac communication,” IEEE Transactions on Biomedical Circuits and Systems, 2024.

[20]  A. Khaleghi, R. Noormohammadi, and I. Balasingham, “Conductive im-  pulse for wireless communication in dual-chamber leadless pacemakers,” IEEE Transactions on Microwave Theory and Techniques, vol. 69, no. 1,  pp. 443–451, 2020.

[21]  M.  Maldari,  M.  Albatat,  J.  Bergsland,  Y.  Haddab,  C.  Jabbour,  and P.  Desgreys,  “Wide  frequency  characterization  of intra-body  commu- nication  for  leadless  pacemakers,”  IEEE  Transactions  on  Biomedical Engineering, vol. 67, no. 11, pp. 3223–3233, 2020.

[22]  L. Chen, Y. Liu, Z. Chen, S. H. Pun, M. I. Vai, and Y. Gao, “An in- vestigation on conductive intracardiac communication dynamic channel gain during the cardiac cycle for leadless pacemakers,” IEEE Journal of

Electromagnetics, RF and Microwaves in Medicine and Biology, vol. 7, no. 1, pp. 82–89, 2022.

[23]  Z. Wei, H. Wang, D. Li, M. I. Vai, S. H. Pun, J. Yang, M. Du, and Y. Gao, “A time-varying equivalent circuit modeling and measuring approach for intracardiac communication in leadless pacemakers,” IEEE Transactions on Biomedical Circuits and Systems, vol. 18, no. 4, pp. 872–884, 2024.

[24]  A. Ryser, T. Schmid, L. Bereuter, J. Burger, T. Reichlin, T. Niederhauser, and A. Haeberlin, “Modulation scheme analysis for low-power leadless pacemaker synchronization based on conductive intracardiac communi- cation,”  IEEE transactions on biomedical circuits and systems, vol.  16, no. 3, pp. 419–429, 2022.

[25]  L. Bereuter, T. Niederhauser, M. Kucera, D. Loosli, I. Steib, M. Schild- knecht, A. Zurbuchen, F. Noti, H. Tanner, T. Reichlin, et al., “Leadless cardiac resynchronization  therapy:  An  in  vivo proof-of-concept  study of wireless pacemaker synchronization,” Heart rhythm, vol.  16, no. 6, pp. 936–942, 2019.

[26]  D. R. Frey, “Improved super-regenerative receiver theory,” IEEE Trans- actions  on  Circuits  and  Systems  I:  Regular  Papers,  vol.  60,  no.  12, pp. 3267–3278, 2013.

[27]  G. MacFarlane and J. Whitehead, “The theory of the super-regenerative receiver  operated  in  the  linear  mode,”   Journal  of  the  Institution  of Electrical Engineers-Part III: Radio and Communication Engineering, vol. 95, no. 35, pp. 143–157, 1948.

[28]  J.-Y. Chen, M. P. Flynn, and J. P. Hayes, “A fully integrated 1.9-ghz auto- calibrated super-regenerative receiver in 0.13-μm cmos,” IEEE Journal of Solid-State Circuits, vol. 42, no. 9, pp. 1976–1985, 2007.

[29]  M.  Vidojkovic,  X.  Huang,  P.  Harpe,  S.  Rampu,  C.  Zhou,  L.  Huang, J. van de Molengraft, K. Imamura, B. Busze, F. Bouwens, et al., “A 2.4 ghz ulp ook single-chip transceiver for healthcare applications,” IEEE transactions on biomedical circuits and systems, vol. 5, no. 6, pp. 523– 534, 2011.

[30]  W. Zhao, D. Li, J. Wang, S. H. Pun, M. I. Vai, and Y. Gao, “An analog front-end  with  automatic  gain  control  for  galvanic-coupled  human body  communication,”  2023  Cross  Strait Radio Science and  Wireless Technology Conference (CSRSWTC), pp. 1–3, 2023.

[31]  Y. Li, L. Yang, L. Yu, B. Liao, and P. Xu, “Digital agc circuit design based on fpga,” Journal of Physics: Conference Series, vol. 1654, 2020.

[32]  D.  Li,  J.  Wang,  H.  Wang,  X.  Huang,  J.  Yang,  Y.  Gao,  H.-C.  Li,  M. I. Vai, and  S. H. Pun,  “A dynamic high-fidelity equivalent circuit  phantom  for  intracardiac  communication  in  pacemaker  indications,” IEEE Transactions on Instrumentation and Measurement, 2025.

[33]  S.  Gabriel,  R.  Lau,  and  C.  Gabriel,  “The  dielectric  properties  of biological tissues: Iii. parametric models for the dielectric spectrum of tissues,” Physics in medicine & biology, vol. 41, no. 11, p. 2271, 1996.

[34]  A. M. Maceira, J. Cosn-Sales, M. Roughton, S. K. Prasad, and D. J.

Pennell, “Reference left atrial dimensions and volumes by steady state free precession cardiovascular magnetic resonance,” Journal of cardio- vascular magnetic resonance, vol. 12, no. 1, p. 65, 2010.

[35]  A.  Maceira,  S.  Prasad,  M.  Khan,  and  D.  Pennell,  “Normalized  left ventricular systolic and diastolic function by steady state free precession cardiovascular magnetic resonance,” J Cardiovasc Magn Reson, vol. 8, no. 3, pp. 417–426, 2006.

[36]  A. M. Maceira, S. K. Prasad, M. Khan, Pennell, and D. J., “Reference right ventricular systolic and diastolic function normalized to age, gender and body surface area from steady-state free precession cardiovascular magnetic resonance,” European Heart Journal, vol. 27, no. 23, pp. 2879– 2888, 2006.

[37]  A. M. Maceira and J. Cosn-Sales. . . , “Reference right atrial dimensions and  volume  estimation by  steady  state  free precession  cardiovascular magnetic resonance,”  Journal of  Cardiovascular Magnetic Resonance, vol. 15, no. 1, pp. 29–29, 2013.

[38]  B. Yuk, B. Kim,  S. Park, Y. Huh, and J. Bae,  “An implantable body channel  communication   system  with   3.7-pj/b  reception   and  34-pj/b transmission  efficiencies,”  IEEE  Solid-State   Circuits  Letters,  vol.  3, pp. 50–53, 2020.